## Slide 1: PCB 24wells
Sofia Peressotti
DSM collaboration

## Slide 3: PCB desgin
PHASE 1.2
Setup and Characterization
Specs:
24 channels
Can be connected to 1 or 2 stimulators (DS5s in this case but it’s flexible)
0-16 mA, 1-
Current and voltage can be measured for each well using an external differential oscilloscope
A filter is applied to remove DC offset (but can be switched on and off) with the capacitors and a 100kohm leakage resistors is there for discharges of the capacitors and potential current leakages

## Slide 4
PHASE 1.2
Setup and Characterization

## Slide 5: EIS results
AUTOLAB LSBI
14.01.26

## Slide 6: The circuit

## Slide 7: All circuit - impedance

## Slide 8: All circuit - phase

## Slide 9: Single arm - Impedance
Working electrode (+)
Counter electrode (-)
Stim1
C=2.2 uF, R2 = 1 ohm
Stim2
C=1 uF, R2 = 10 ohm

## Slide 10: Single arm - Phase
Working electrode (+)
Counter electrode (-)
Stim1
C=2.2 uF, R2 = 1 ohm
Stim2
C=1 uF, R2 = 10 ohm

## Slide 11: WE vs CE
Stim1
Stim2
Stim1
C=2.2 uF, R2 = 1 ohm
Stim2
C=1 uF, R2 = 10 ohm

## Slide 12: Values @ 1 kHz
Working electrode (+)
Counter electrode (-)
All circuit
Stim1
R2 = 1 ohm
Stim2
R2 = 10 ohm
With (actual)
filter OFF
Well impedance:
Stim1: 158.6 ohm
Stim2: 159.8 ohm
With (actual)
filter ON
Well impedance:
Stim1: 113.1 ohm
Stim2: 81.0 ohm

## Slide 13: Values @ 10 Hz
Working electrode (+)
Counter electrode (-)
All circuit
Stim1
R2 = 1 ohm
Stim2
R2 = 10 ohm
With (actual)
filter OFF
Well impedance:
Stim1: 158.6 ohm
Stim2: 159.8 ohm
With (actual)
filter ON
Well impedance:
Stim1: 113.1 ohm
Stim2: 81.0 ohm

## Slide 14: Fit and simulation(Autolab) – all circuit, well: A1
Stim1_0 – A1
Stim1_1 – A1
Stim2_0 – A1
Stim2_1 – A1

## Slide 15: Fit and simulation(Autolab) – WE arm, well: A1
Stim1_0 – A1
Stim1_1 – A1
Stim2_0 – A1
Stim2_1 – A1

## Slide 16: Fit and simulation(Autolab) – CE arm, well: A1
Stim1_0 – A1
Stim1_1 – A1
Stim2_0 – A1
Stim2_1 – A1

## Slide 17: Circle fit (Autolab) – all circuit, well: A1 FIXED
Stim1_0 – A1
Stim1_1 – A1
Stim2_0 – A1
Stim2_1 – A1
Rcomponent:
24.0x4=96
25.9x4=103.6
26.6x4=106.4
27.2x4=108.8
Avg=103.7 ohm
From single arm: 21.4x4=85.6 ohm
MUX: 50ca.x2
IC6: 5 ca. x2
Expected: 110 ohm

## Slide 18: Circle fit (Autolab) – all circuit, well: A1
Stim1_0 – A1
Stim1_1 – A1
Stim2_0 – A1
Stim2_1 – A1

## Slide 19: Values @ 1 kHz circuit - well

## Slide 20: Circle fit (Autolab) – well PEDOT vs Pt
D1_Pt
D2_Pt
D3_Pt
A1_PEDOT
A3_PEDOT
A6_PEDOT
Pt Average
Rs = 161 ohm
Cdl = 224.3 uF
Rct = 4.45 kohm
PEDOT Average
Rs = 163 ohm
Cdl = 653.5 uF
Rct = 1.58 kohm
NB: A1 was excluded because the fitting was not accurate

## Slide 21: 24device - well only

## Slide 22: Simulation – Stim1_0

## Slide 23: Simulation – Stim1_1

## Slide 24: Conclusions
The impedance of the well only at 1 kHz is 160 ohm ca. (at 1mA, expected EF is )
impedance of the circuit is around 300 ohm at >1000 Hz and without capacitors
Capacitors make the impedance at low frequency too high to be used, so the filter will only be activated at >1000 Hz
The capacitors shift the phase quite a lot
Stim 2 had a lower impedance of the circuit
Rmeasure resistor is ok at 1 ohm to minimize difference between WE and CE
PEDOT and Pt have comparable impedance and behaviour. Impedance only changes a lot at 1-5 Hz. I will probably keep it Pt to minimize variability
I can calculate the actual impedance of the well for each frequency and estimate the well Voltage, assuming circuit impedance is 270 ohm and the well impedance is 160 ca.
The impedance and phase values at 1kHz are reflected in the simulations
To calculate the impedance of the well, I will remove the average impedance of the circuit at the given frequency based on average ON and OFF values. The phase will also be recorded
Next:
change both MUX to stim 2 conditions
Remove PEDOT and measure all circuit values again
Validate the well voltage values with oscilloscope
Build GUI to operate PCB and to visualize the voltage in the well real time and record expected impedance and phase shift
Electrode-electrolyte interface values are:
PEDOT Average
Rs = 163 ohm
Cdl = 653.5 uF
Rct = 1.58 kohm
Pt Average
Rs = 161 ohm
Cdl = 224.3 uF
Rct = 4.45 kohm